## Comments
| Report: |  | Differential Pressure Transmitter Calibration Optimization
| Project: |  | MEAS-2-22
| Contractor |  | Probaren Institute of Technology
| Contract Number |  | PR027-14201
| Reviewer comments | [5 empty] | Contractor response
Item | By | Document Version | Original Page, Paragraph | Section Type | Type of Comment | Comment | By | Revision Version | New Page, Paragraph | Revision Comment
1 | Choquette | Draft 1 | 1 |  | Conformance | Cover page does not conform to PRCI requirements | Smith | Draft 2 | 1 | Cover page updated
2 | Choquette | Draft 1 | Multiple |  | Conformance | Catalog number should be PR-027-14201-R01 | Smith | Draft 2 | Multiple | Updated catalog number
3 | Choquette | Draft 1 | 3 |  | Conformance | Missing a 'value to member section' | Smith | Draft 2 | 3 | Added a value to member section
4 | Becket | Draft 1 | 6, 2 | Section/Clause | Technical | The audit trail should include the original data and edits. If this is implied, it is not clear. | Jones | Draft 2 | 6, 3 | Added a statement "Original data must be retained and part of the audit package."
5 | Benton | Draft 1 | Multiple | All Sections | Editorial | Coriolis is a proper name it should be capitalized throughout the document | Smith | Draft 2 | Multiple | Capitalized Coriolis throughout the document
6 | Gibson | Draft 1 | 6,5 | All Sections | Technical | The new approach to define the averaging method for DP, Ps, and Tf as flow dependent linear average in various parts of 1.4 is in direct opposition to Equation 6 which allows flow dependent formulaic averaging under certain conditions. | Jones | Draft 2 | 7,4 & Appendix C | Deleted "linear" from all locations that are not specific to determining DPdynamic to include differential pressure, static pressure, and flowing temperature. The reader is also directed to a new Appendix C for a discussion on the acceptable averaging methods.
7 | Holbrook | Draft 1 | 34, 3 | Intro | Editorial | Hard to understand the last sentence | Jones | Draft 2 | 35, 2 | Added a simple example to provide more explanation.
8 | Benton | Draft 1 | 9 & 10 | Section/Clause | General | For Figure 8 and Figure 9, ISA symbols are being used. Include a reference in section 4.3 | Jones | Draft 3 | 13, 1 | A new section has been added in 4.3 that includes a cross reference to ISA symbols.
9 | Benton | Draft 2 | 3, 4 | All Sections | Editorial | Coriolis is not capitalized | Jones | Draft 3 | 4,1 | Capitalized Coriolis
10 | Jackson | Draft 2 | 8, 2 | Section/Clause | Editorial | What are "other approved linear metering standards"? Too vague, needs to be tighten up or deleted. | Jones | Draft 3 | 8, 2 | Deleted phrase "or other approved linear metering standards."